README.DOC Version 2 (June 1995)

Tektronix DocuWave Version 1.2

(c) Copyright 1995 Tektronix, Inc.

Portions (c) Copyright 1995 National Instruments, Inc.

All rights reserved.

Contents:

This document contains the following sections:

DocuWave Version 1.2 System Requirements

Windows 3.x Support

Excel 4.0 Support

Supported Tektronix Oscilloscopes

RS232 Basic Information

RS232 Operation

-11000 Series

-TDS 300 Series

-TDS 400/500/600/800 Series

-2221A

-THM 500 Series & Derivatives (ET2020, AVL865, etc.)

-THS 700 & THS DMM Series & Derivatives

RS232 Hints

RS232 Cable Issues

RS232 Cable Examples

-11000 Series to PC (DB25 to DB9)

-11000 Series to PC (DB25 to DB25)

-TDS 300 Series To PC

-DB25 to DB25 Null Modem

Existing Driver Modification Examples

-TDS400/A driver modification

-11800 and CSA RS-232 operation

-11000 series waveform quantity limit

-2221 and 2221A RS-232 operation

Adding New Drivers

Commonly Asked Questions

Application-specific Notes

To install DocuWave on your system from the distribution disk,

change your directory to the drive containing the disk and

type "setup".

DocuWave Version 1.2 System Requirements:

=============================================================

DocuWave has the following system requirements:

CPU:         286 or better

Memory:      520k for full functionality DocuWave may work

with less but the TIFF, PCX, and BMP image

creation operations may run out of memory.

DocuWave may not work with some networks

(the remaining amount of memory after loading the

network drivers is generally less than the

required 520K for full functionality).

Disk Space:  1.5 Megabytes

MS DOS:      3.x and above.

DocuWave supports both RS-232 and GPIB interfaces.

The GPIB interface must be IEEE 488.2 compatible.

The following National Instruments GPIB products are

compatible with DocuWave:

Hardware/Software Kits:

AT-GPIB/TNT          P/N 776786-01

GPIB-PCII            P/N 776092-01

GPIB-PCIIA           P/N 776113-01

MC-GPIB              P/N 776183-01

PCMCIA-GPIB          P/N 776857-01

Windows 3.x Support:

=============================================================

DocuWave Version 1.2 may be run as a full screen DOS application

under Windows with the following exceptions:

When used with Windows 3.x, 32 bit disk access should not be used.

Memory before entering Windows should be at least 520k.

DocuWave may not work with some networks

(the remaining amount of memory after loading the

network drivers is generally less than the

required 520K for full functionality).

After installing DocuWave on your system, you can use

"wininstl.exe" to aid in starting DocuWave from windows.

To do this, start windows, select Run... from the File menu, and

specify "wininstl.exe" which exists in the same directory as the

DocuWave executable.  This will create a DocuWave program group

and the DocuWave run and setup items.

Excel 4.0 Support:

=============================================================

In the same directory as the DocuWave executable, the following

template files have been added to aid users of Excel 4.0 with

importing waveforms that have been exported from DocuWave:

docuwave.xla

docuwave.xlm

docuwave.xls

Supported Tektronix Oscilloscopes:

=============================================================

11K/DSA          GPIB         RS-232

-----------        -----        ------

Tek 11401            Y            Y

Tek 11402            Y            Y

Tek 11403/A          Y            Y

Tek 11801/B          Y            Y

Tek 11802            Y            Y

Tek CSA803           Y            Y

Tek DSA601/A         Y            Y

Tek DSA602/A         Y            Y

2200/2400         GPIB         RS-232

-----------        -----        ------

Tek 2212           Y w/OPT 10   Y w/OPT 12

Tek 2216           Y w/OPT 10     N

Tek 2221/A         Y w/OPT 10   Y w/OPT 12

Tek 2230           Y w/OPT 10   Y w/OPT 12

Tek 2232           Y w/OPT 10   Y w/OPT 12

Tek 2430/A           Y            N

Tek 2432             Y            N

Tek 2440             Y            N

RTD/SCD          GPIB         RS-232

------------       -----        ------

Tek RTD720/A         Y            N

Tek SCD1000          Y            N

Tek SCD5000          Y            N

TDS            GPIB         RS-232

------------       -----        ------

Tek TDS310         Y w/OPT 14   Y w/OPT 14

Tek TDS320         Y w/OPT 14   Y w/OPT 14

Tek TDS350         Y w/OPT 14   Y w/OPT 14

Tek TDS410/A         Y          N (Printer Port Only)

Tek TDS420/A         Y          N (Printer Port Only)

Tek TDS460/A         Y          N (Printer Port Only)

Tek TDS520/A         Y          N (Printer Port Only)

Tek TDS524/A         Y          N (Printer Port Only)

Tek TDS540/A         Y          N (Printer Port Only)

Tek TDS544/A         Y          N (Printer Port Only)

Tek TDS620/A         Y          N (Printer Port Only)

Tek TDS640/A         Y          N (Printer Port Only)

Tek TDS644/A         Y          N (Printer Port Only)

Tek TDS684/A         Y          N (Printer Port Only)

Tek TDS744/A         Y          N (Printer Port Only)

Tek TDS784A          Y          N (Printer Port Only)

Tek TDS820           Y          N (Printer Port Only)

THM            GPIB         RS-232

-----------        -----        ------

Tek THM550           N            Y w/THMCOM1

Tek THM560           N            Y w/THMCOM1

Tek THM565           N            Y w/THMCOM1

Tek THM571 (ET2020)  N            Y w/THMCOM1

Tek THM575 (AVL865)  N            Y w/THMCOM1

THS            GPIB         RS-232

-----------        -----        ------

Tek THS710           N            Y

Tek THS720           N            Y

Tek THS7DMM          N            Y

RS-232 Basic Information:

=============================================================

Make sure that the correct RS-232 cable is being used.

There are two types of RS-232 devices: Data Terminal Equipment

(DTE) and Data Communications Equipment (DCE).  Most printers,

plotters, and PCs are DTE devices.  Most modems are DCE devices.

The type becomes important when interfacing between devices.

Use a straight-through RS-232 cable when connecting

different types (DTE to DCE) together.

Use a null-modem cable when connecting similar types

(both DTE or both DCE).  A null-modem cable lets similar

devices communicate by crossing the send and receive pairs

and crossing the handshake pairs.

RS-232 Operation:

=============================================================

These are instructions for specific instruments.  If following

these examples does not help, proceed to RS-232 Hints and Driver

Modifications.

--11000 Series

If you're using an 11000 series scope (includes

DSA and CSA series), make sure the following are

set in the Utility/RS232 menu:

- Echo must be Off.

- Verbose must be Off.

- Debug must be Off.

- Delay should be 0.

If you're using an 11000 series scope (includes

DSA and CSA series), you can use the "Debug"

facility to see if any characters are being sent

from DocuWave to the scope. To use this facility,

go to the Utility/RS232 menu and toggle the

Debug selector to "On". The top 4 lines of the

instrument screen will display any characters

which come in over the RS232 port.

Be sure and set the Debug value to "Off" when

you're done. DocuWave data transfers will *not*

work when Debug is On.

--TDS 300 Series

If you're using a TDS 300 series scope, you can

use the "Set RS232 Parameters To Default Values"

button which sets the RS232 settings on the

scope to the same as the DocuWave default values.

The button is found in the RS-232C menu. Get to it

by pushing the UTILITY button. Then continually

push the bottom left bezel button until "I/O" is

selected. Then push the bottom bezel button

labeled "RS-232C". The side bezel buttons now

control RS232 settings. The fourth side bezel

button will set the RS232 settings to default

values.

If you're using a TDS 300 series scope, you can

examine the Error Log for clues. The Error Log

is accessed via the Error Log menu. Push Display

Log in that menu to see the Error Log. Examine

the lines entitled "RS232 Line Snapshot" and

"RS232 Errors". See your TDS 300 Programmer

Manual for details on how to interpret the

information in the Error Log.

Get to the Error Log menu by pushing the UTILITY

button. Then continually push the bottom left

bezel button until "Diag" is selected.  Then push

the bottom bezel button labeled "Error Log".

Then push the side bezel button labeled

"Display Log".

--TDS 400/500/600/800 series

The TDS 400/500/600/800 series of scopes support

RS-232 printers but cannot be controlled over

RS-232. DocuWave can only communicate with these

particular scopes using GPIB.

--2221A

The following hints apply to the 2221A. Other

2200 series scopes may or may not work the same

way.

- Make sure the oscilloscope is in "STORE" mode

and not "NON-STORE" as selected by the front

panel button in the upper right hand corner.

- After changing the dip switch settings on the

scope, you must cycle the power in order for the

new settings to take affect.

- Use a baud rate of 2400 or less. The instrument

actually has better performance at 2400 baud than

at faster baud rates.

(Switches 1 thru 4 = 0101 for a 2400 baud rate.)

- Set the parity switch to off (switch 5 = 0)

- Set the EOL switch to CR/LF (switch 8 = 1) and

set the DocuWave terminator to CR-LF in the

Custom RS232 Configuration menu.

- Switches 6, 7, 9, and 10 are don't care.

- In DocuWave, set Data Bits to 8, Stop Bits to 1,

Parity to None, and make sure software handshaking

is *not* selected.

- In the 2221 Communications Menu, set Stop Bits

to One and Flow to OFF.

- Make sure you are using the correct port (DCE

or DTE) for your cable. The sex of the cable is

not necessarily an indication of which port you

should be using.

--THM 500 Series & Derivatives (ET2020, AVL865, etc.)

- THM series RS232 communication parameters are

fixed at 8 bit data, 1 stop bit, no parity, with

selectable 1200 or 9600 baud.  There is neither

hardware nor software flagging.  EOL is CR.

Make sure that DocuWave and the instrument are

both set to the same parameter values.

- In order to minimize communication line time-out

problems, it is recommended that DocuWave be set

for 100 seconds (maximum).

- Since there is no flagging in effect, certain

operations may fail or appear to "lose characters"

when operating at 9600 baud.  If this occurs,

change both DocuWave and the instrument to 1200

baud and repeat the operation.

RS232 Hints:

=============================================================

1.  Make sure that the communication parameters are set

correctly on *both* the instrument and DocuWave as

follows:

- Baud rates must be the same.

- Parity should be none.

- Data width should be 8.

- Stop bits should be 1.

Soft flagging should never be used for waveform and

hardcopy transfers.  This is because waveform data and

hardcopy output contain binary characters which can

easily include the XON and XOFF characters used to

implement soft flagging. Soft flagging should only be

used for transferring stored setups since stored setup

information generally consists of printable ASCII

characters.

2.  It generally shouldn't matter if the EOL (end of line)

characters are the same but just to be sure, the

recommended setting is CRLF.

3.  Normally DocuWave "remembers" the settings which were

selected the last time DocuWave was used. To force

DocuWave to use the default settings, exit DocuWave and

delete the file called DOCUWAVE.PRO.

4.  Not all instruments support full RS232 communications.

Some instruments only support GPIB communications. For

some instruments, RS232 is an option that may or may not

be installed. Other instruments may only support a

limited RS232 capability (ie. printer functions only).

5.  If you are able to connect to an instrument but unable

to transfer waveforms successfully, you may have a cable

which is not wired correctly for hard flagging. Try the

following to verify this:

a. Can you successfully transfer settings to and from

the instrument at the fastest supported baud rate? If so,

hard flagging is probably working. If not, try soft

flagging.

If you can successfully transfer settings using soft

flagging but not with hard flagging, there is a good

chance the hard flagging lines of your cable are not

connected properly.

Make sure that the *entire* setting is restored when you

send the setting back to the instrument. It is easy for

a setting to be *partially* restored if hard flagging is

not working or if you are not using any flagging.

b. Try using no flagging of any kind and use a slow baud

rate. If waveform transfers work under these conditions,

there is a good chance the hard flagging lines of your

cable are not connected properly.

RS232 Cable Issues:

=============================================================

If all of the above fails to solve your problem, it may be that

you are using an incompatible RS232 cable. Unfortunately, there

is no such thing as a standard cable. You may need to use an

ohmmeter to verify that your cable is wired correctly. Many

cables don't have all the pins connected.

The first thing you need to know is whether the PC and

instrument are DTE (Data Terminal Equipment) or DCE (Data

Communications Equipment) devices. This information is in the

documentation which comes with the instrument and PC.

If you are connecting a DTE device to a DCE device, you need a

"straight-through" cable. For example, the transmit pins of both

the instrument and PC should be connected together.

If you are connecting two devices of the same type, you need a

"null-modem" cable. This is a cable whose lines are crossed such

that the inputs of one device are connected to the outputs of

the other. For example, the transmit pin of the instrument

should be connected to the receive pin of the PC and vice versa.

WARNING! Do not connect the output lines of two same-type

devices to each other. Doing so may damage one or both of the

devices.

For hard flagging to work, the RTS (Request to Send) and CTS

(Clear to Send) lines must be connected. Use an ohmmeter to

verify that they are actually connected since not all cables

connect these lines.

The DCD (Data Carrier Detect) line is not used by DocuWave so it

doesn't matter whether or not it is connected. The DCD line is

sometimes referred to as RLSD.

The RTS (Request to Send) line is sometimes referred to as RFR.

RS232 Cable Examples:

=============================================================

11000 Series To PC:

-------------------

--11000 Series To PC (DB25 male to DB9 female.)

The following represents the correct wiring for connecting

an 11000 series (including DSA and CSA series) instrument with

25 pins to a 9 pin serial port on a DTE personal computer.

Note that 11000 series instruments are DCE devices.

Example: Tektronix part number 012-1241-00 (15 feet)

DB25 Male                     DB9 Female

11K Line              Pin Number  PC Pin           Line Name

--------------------------------  ---------------------------

Chassis Ground            1       Connector Shell

Transmit Data (TxD)       2       3                  TxD

Receive Data (RxD)        3       2                  RxD

Request to Send (RTS)     4       7                  RTS

Clear To Send (CTS)       5       8                  CTS

Data Set Ready (DSR)      6       6                  DSR

Signal Ground (GND)       7       5                  GND

Data Carrier Detect (DCD) 8       1                  DCD

Data Terminal Ready (DTR) 20      4                  DTR

--11000 Series To PC (DB25 female to DB25 male.)

The 11000 series instruments are DCE devices.

An example cable is the Tektronix 012-0911-00.

DB25 Female                    DB25 Male

11000 Line            Pin Number  PC Pin            Line Name

--------------------------------  ---------------------------

Ground                    1       1                  GND

Transmit Data (TxD)       2       2                  RxD

Receive Data (RxD)        3       3                  TxD

Request to Send (RTS)     4       4                  CTS

Clear To Send (CTS        5       5                  RTS

Data Set Ready (DSR)      6       6                  DTR

Signal Ground (GND)       7       7                  GND

Data Carrier Detect (DCD) 8       20                 DCD

Data Terminal Ready (DTR) 20      8                  DSR

--TDS 300 Series To PC  (DB9 female to DB9 female)

The following represents the correct wiring for connecting a

TDS 300 series oscilloscope to a 9 pin serial port on a DTE

personal computer.  Note that TDS 300 series scopes are DTE

devices.

Example:  Tektronix part number for such a cable is 012-1379-00.

This cable has a female connector on each end and is

six feet long.

DB9 Female                     DB9 Female

TDS Line              Pin Number  PC Pin            Line Name

--------------------------------  ---------------------------

Receive Data (RxD)        2       3                  TxD

Transmit Data (TxD)       3       2                  RxD

Data Terminal Ready (DTR) 4       6                  DSR

Signal Ground (GND)       5       5                  GND

Data Set Ready (DSR)      6       4                  DTR

Request to Send (RTS)     7       1, 8               CTS

Clear To Send (CTS)       8       7                  RTS

"Note that pins 1 and 8 are wired together."

--DB25 to DB25 Null Modem

This is the cable type to use when joining two similar devices

together. DCE to DCE.  DTE to DTE.

Example:  DB25 female to DB25 female 012-1423-00  (5' cable)

DB25 male to DB25 male 131-4923-00 (adapter)

DB25 Female                    DB25 Female

25 Pin Connector      Pin Number  PC Pin           Line Name

--------------------------------  ---------------------------

Chassis Ground         1            1       Chassis Ground

TxD                    2            3                  RxD

RxD                    3            2                  TxD

RTS                    4            5                  CTS

CTS                    5            4                  RTS

DSR DCD                6,8          20                 DTR

DTR                    20           6,8            DSR DCD

Signal GND             7            7           Signal GND

EXISTING DRIVER MODIFICATIONS EXAMPLES

=============================================================

The follow examples show how to modify DocuWave's Scope Drivers

and Driver lists.  This section is also describes fixes and

work arounds to some instruments.  The listed changes may be required

for older instrument software and firmware.

Existing Driver Change Examples:

--TDS Series (TDS410A GPIB only)

No TDS410A driver listed in the program.

Description:

The drivers list does not include the scope I am using.

Driver Modifications/Creation:

To add an oscilloscope to an existing driver follow

these steps.

First)  Modify the drivers.lst file

Second) Modify the closest scope.dvr file

Notes:

Before modifying any files, it is a good idea to make a

backup copy of the original. It is recommended that you

modify the DocuWave files which have been installed on

your hard drive rather than modifying the files on the

actual installation disk.

Procedure:

1.  Change to the DOCUWAVE\DRIVERS directory on your PC.

2.  Edit the file DRIVERS.LST using a text editor.

Add the TDS410 and TDS410A to the DRIVERS LIST as

indicated:

DRIVERS.LST Before:

3

Tek TDS350:TDS300.DVR:.350

Tek TDS420/A:TDS400.DVR:.420

Tek TDS460/A:TDS400.DVR:.460

DRIVERS.LST After:

4                            <this line is the number of instruments

Tek TDS350:TDS300.DVR:.350

Tek TDS410/A:TDS400.DVR:.410 <this line has been added

Tek TDS420/A:TDS400.DVR:.420

Tek TDS460/A:TDS400.DVR:.460

At the top of the "drivers.lst" file is a number

which must reflect the total number of instruments

included in the file (i.e., the number of lines).  In

the example above the number needs to be modified

from "3" to "4" in this case.

The DRIVERS.LST file is now modified.

3.  Next, modify the actual scope driver.

Copy the existing scope.dvr file to create a backup

file.  Enter the DOS command: COPY TDS400.DVR TDS400.OLD

4.  Edit the file TDS400.DVR.

4a. Change the second line as indicated:

Before:

TDS 420`TDS 460

After:

TDS 410`TDS 420`TDS 460

5.  Change to the DOCUWAVE directory and type:

DOCUWAVE

Select the appropriate instrument. When the connection

is successful you should see the string "TDS410" in the

upper right hand corner.

--11800 and CSA Oscilloscopes (RS-232 operation)

RS232 doesn't work with 11800 and CSA series oscilloscopes.

Description:

Although you can *connect* to an 11800 or CSA series

scope, you will be unable to make any waveform, setting,

or hardcopy transfers. DocuWave will either report a

time-out error or other error when you attempt to do so.

Driver Modifications

Work around:

The following steps will enable DocuWave to work

correctly with an 11800 and/or CSA series oscilloscope.

This work around involves making modifications to the

DocuWave driver files and as such, it is a one-time

modification.

Notes:

Before modifying any files, it is a good idea to make a

backup copy of the original. It is recommended that you

modify the DocuWave files which have been installed on

your hard drive rather than modifying the files on the

actual installation disk.

Procedure:

1.  Change to the DOCUWAVE\DRIVERS directory on your PC.

2.  Edit the file DRIVERS.LST

Change the string "11K" to "11800" in the following

three lines as indicated:

Before:

Tek 11801:11K.DVR:.801

Tek 11802:11K.DVR:.802

Tek CSA803:11K.DVR:.803

After:

Tek 11801:11800.DVR:.801

Tek 11802:11800.DVR:.802

Tek CSA803:11800.DVR:.803

3.  Enter the DOS command: COPY 11K.DVR 11800.DVR

4.  Edit the file 11800.DVR which you just created.

4a. Change the first line as indicated:

Before:

3.0`ID?

After:

3.0`DATE?

4b. Change the second line as indicated:

Before:

11401`11402`11403`11801`11802`CSA`DSA601`DSA602

After:

DATE

5.  Change to the DOCUWAVE directory and type:

DOCUWAVE

Select the appropriate instrument. When the connection

is successful you should see the string "DATE" in the

upper right hand corner. Normally you would see the

instrument name in that location but this is a caveat

of this work around.

--11000 Series Oscilloscopes

There are only 8 stored waveforms available for 11000 series

scopes.

Description:

This is an arbitrary limitation which you can change.

However, the DocuWave menus are limited to 30 entries

which means if you have 8 traces in the menu, you can

only have 22 stored waveforms. The 22 stored waveforms

don't have to be the first 22 however. They can be any

22 you like.

Notes:

Before modifying any files, it is a good idea to make a

backup copy of the original. It is recommended that you

modify the DocuWave files which have been installed on

your hard drive rather than modifying the files on the

actual installation disk.

Procedure:

1.  Change to the DOCUWAVE\DRIVERS directory on your PC.

2.  Edit the file 11K.DVR

Modify lines 5 and 6 as desired. Line 5 contains the

text which appears in the DocuWave menu and line 6

contains a corresponding command name for each line in

the menu. There must be a one for one match between the

line 5 items and the line 6 items.

Example:

Line 5: Trace 1`Trace 2`Store 125`Store 126

Line 6: OUT `TRA1`TRA2`STO125`STO126

Besides the limit of 30 menu entries, there is a limit

of 512 characters on any given line of the .DVR file.

--2221 and 2221A Oscilloscopes

RS232 doesn't work with the 2221 or 2221A oscilloscopes.

Description:

Although you can *connect* to an 2221 scope, you will be

unable to make any waveform or hardcopy transfers.

DocuWave will either report a time-out error or other

error when you attempt to do so.

Work around:

The following steps will enable DocuWave to work

correctly with a 2221 or 2221A oscilloscope. This

work around involves making modifications to the DocuWave

driver files and is a one-time modification.

Notes:

Before modifying any files, it is a good idea to make a

backup copy of the original. It is recommended that you

modify the DocuWave files which have been installed on

your hard drive rather than modifying the files on the

actual installation disk.

Procedure:

1.  Change to the DOCUWAVE\DRIVERS directory on your PC.

2.  Edit the file 2221.DVR

Change the first character on line 8 from a '0' to a

'1' as indicated:

Before:

0`LONG OFF;DAT ENC:BIN

After:

1`LONG OFF;DAT ENC:BIN

3.  Change back to the DOCUWAVE directory and type 'DOCUWAVE'

and verify that you can now make waveform transfers. If it

still doesn't work, refer to the 2221A hint in the RS232

Hints section.

ADDING NEW INSTRUMENT DRIVERS TO DOCUWAVE

=============================================================

The follow examples show how to modify DocuWave's Scope Drivers

and Driver lists.  This section is also describes fixes and

work arounds to some instruments.  The listed changes may be required

for older instrument software and firmware.

For an instrument to work with DocuWave, two actions must happen:

1)  Create an instrument driver file and place it within

the DocuWave drivers directory.  (Details Below)

2)  Update the file "drivers.lst" in the DocuWave drivers

directory.  (Details Below)

If the current drivers.lst file looks like this:

6

Tek 11401:11K.DVR:.401

Tek SCD1000:SCD.DVR:.SC1

Tek SCD5000:SCD.DVR:.SC5

Tek TDS320:TDS300.DVR:.320

Tek TDS420:TDS400.DVR:.420

Tek TDS460:TDS400.DVR:.460

and for example, we need to add a TDS410, then make the following

changes

7                             <this line reflects the number of instruments

Tek 11401:11K.DVR:.401

Tek SCD1000:SCD.DVR:.SC1

Tek SCD5000:SCD.DVR:.SC5

Tek TDS320:TDS300.DVR:.320

Tek TDS410:TDS400.DVR:.410    <this line has been added

Tek TDS420:TDS400.DVR:.420

Tek TDS460:TDS400.DVR:.460

The added line consists of 3 fields separated by colons:

1)  The name of the instrument as it is to appear in the

selection list.

2)  The name of the driver file for the instrument.  This

file must be in the drivers directory of DocuWave.

3)  The extension ( including the .) to be used by DocuWave

when saving files.

DocuWave 1.0 Instrument Driver File Format Information

-------------------------------------------------------

The DocuWave instrument drivers files are ASCII text files containing

up to 18 lines of data plus any number of comment lines.  Lines

containing comments begin with a forward slash character (/) and are

ignored by DocuWave.  The lines containing data are broken into fields

by the back single quote character (`).  Each line has a limit of 255

characters.

LINE 1:  2 fields

1)  Number representing the version of the driver file.

2)  Command to query the instrument for its ID.

Sample:  3.0`*IDN?

LINE 2:  1 or more fields

1)  String representing the name of the instrument that may

be returned by the instrument after sending the query in

LINE 1.  (Spaces in the string are ignored for comparison)

2, 3, ...)  Same as 1

Sample:  TDS 320`TDS 350

LINE 3:  1 field

1)  String to uniquely identify the family

Sample:  TDS

LINE 4:  1 or 2 fields

1)  Either a zero or non-zero value.

Zero means the instrument cannot be queried for its

settings and field 2 does not exist.  Otherwise field

2 must exist.

2)  Command to query instrument settings.

Sample:  1`*LRN?

LINE 5:  1 or more fields

1)  Readable name for a waveform source on the instrument.

2, 3, ...)  Same as 1.

Sample:  Ch 1`Ch 2`Ref 1

LINE 6:  2 or more fields (dependent on LINE 5)

1)  Source selection command prefix.

This is the command prefix to select a source on the instrument

2)  Source selection command suffix.

This is the command suffix corresponding to the source name

in LINE 5 field 1.

3, 4, ...)  Same as 2 but corresponds to successive fields in LINE 5.

LINE 7:  1 or more fields

1)  Either a zero or non-zero value.

Zero means the instrument does not have records and no other

fields exist.

Otherwise fields 2, 3, and 4 must exist.

2)  Record selection command prefix.

This is the command prefix to select a record on the instrument.

3)  Value to be used as the default record.

4)  Reserved, any value may be present.

Sample:  1`DATA STRE:`1`1

LINE 8:  2 fields

1)  Either a zero or non-zero value.

Zero means the instrument does not respond to setup commands

sent over RS232.

Otherwise the instrument responds with some type of "READY"

string.

2) Setup command to initialize the instrument to a known state.

This command can have an effect on other lines:  setting the

number of bytes per point, for example, would have to be

reflected in LINE 10, field 3.

Sample:  HEAD ON;:VERB OFF;:DATA:ENC RIB;WID 1

LINE 9:  8 fields

1)  Command to query the instrument for the waveform preamble.

2)  Search string used to search the preamble for the point format.

3)  Search string used to search the preamble for the y-multiplier.

4)  Search string used to search the preamble for the y-offset.

5)  Search string used to search the preamble for the y-zero.

6)  Search string used to search the preamble for the x-increment.

7)  Search string used to search the preamble for the x-offset.

8)  Search string used to search the preamble for the x-zero.

If any of the strings in fields 2 through 8 contain a tilde (~),

the tilde will be replaced by the index of the current source

specified in LINE 5 prior to the search.

For Example:  YZE:CH~: may become YZE:CH2:

Sample:  WFMP?`PT_F`YMU`MOF`XIN`PT_O`

LINE 10: 3 fields

1)  Character specifying the waveform file type returned by the

instrument.

#  IEEE 488.2 format

%  Tektronix format

2)  Command to have the instrument send a waveform.

3)  Number of bytes per waveform point.

If negative, it indicates that a checksum is sent with IEEE

488.2 format.

4)  A numeric value specifying additional information about

the waveform.

If field 3 is 1 (or -1), a zero means data is sent as signed

bytes, otherwise they are unsigned.

If field 3 is 2 (or -2), then a 2 here indicates the data is

sent MSB first.

5)  Reserved for special processing for the RTD and SCD series.

Sample:  #`CURV?`1`0

LINE 11: 3 fields.

1)  Number of digitizing levels per vertical division.

2)  Digitized value at center screen.

3)  Number of vertical divisions.

Sample:  25`0`8

LINES 12 - 15:

These lines determine whether the instrument can have waveforms

sent back to it.  If LINE 12 is not blank, then LINES 13, 14 and

15 must not be blank and they indicate how to send a waveform

back to the instrument.

LINE 12: 0 or more fields.

1)  Readable name for a waveform destination on the instrument.

2, 3, ...)  Same as 1.

Sample:  Ref 1`Ref 2`Ref 3

LINE 13: 0 or more fields (dependent on LINE 12)

1)  Destination selection command prefix.

This is the command prefix to select a destination on the

instrument.

2)  Source selection command suffix.

This is the command suffix corresponding to the source

name in LINE 12, field 1.

3, 4, ...)  Same as 2 but corresponds to successive fields in

LINE 12.

Sample:  DATA:ENC RIB:START 1:`TAR`REF1`REF2`REF3

LINE 14: 0 or more fields.

1, 2, ...)  Waveform preamble prototype.

Each field is part of the entire preamble.

The preamble is assembled by placing the values

specified by LINE 15 at the locations where the

fields are divided.

Sample:  DELE:WAVEFORM`;:ALLO:WAVEFORM:``;:DATA:TAR`;

LINE 15: 0 or more fields (dependent on LINE 14)

1, 2, ...)  Waveform preamble variable specifiers.

The specifiers indicate what values to place in

the preamble prototype in LINE 14.  The first

specifier matches the second field marker in

LINE 14, and so on.

Valid specifiers are:

Target  - replaced with index of destination

on LINE 13.

#       - replaced with number of points

PTFMT   - replaced with point format

XZERO   - replaced with x-zero

XINC    - replaced with x-increment

XOFF    - replaced with x-offset

YZERO   - replaced with y-zero

YMULT   - replaced with y-multiplier

YOFF    - replaced with y-offset

Sample:  TARGET`TARGET`#`TARGET

LINE 16: 0 or 1 field.

1)  If the instrument has a hardcopy feature, this field

contains the command to start it.

Sample:  :HARDC START;

LINE 17: 0 or 5 fields.

1)  Command to setup instrument error checking.

2)  Command to query if any errors to be reported.

3)  Value returned by command in field 2 if there are

no errors.

4)  Command to get the next error.

5)  Value returned by command in field 4 if there are

no more errors.

Sample:  DESE 56`*ESR?`0`EVMSG?`0

LINE 18: 0 or 1 field.

1)  Anything in this field indicates the instrument responds

to an RS232 com-break.

Sample:  DCL on ComBreak

"This is the end of the "Driver File Format Section"

Commonly Asked Questions:

=============================================================

Q:  Why is the horizontal scale factor on the DocuWave screen

not the same as that on the scope?

A:  Some scopes (11000 series, DSA 600 series, CSA series, 2200

series) use record lengths that are powers of two. Other scopes

(TDS series) use record lengths that are a factor of 10. You'll

notice on an 11K scope that a 512 point record actually spills

over the graticule area with 6 points appearing to the left of

the graticule and 5 points to the right.

DocuWave was designed to display the entire waveform in a

10 graticule display area and so the DocuWave time per div

reflects this.

In other words, the DocuWave time per div reflects the DocuWave

graticule which is not necessarily the same as the scope

graticule.  Try the DocuWave zoom function and you'll see that

the scale factors change to reflect the new graticule.

Version 1.1 and greater:  When using scopes with longer

record lengths the displayed horizontal scale factor will appear

to be incorrect by a factor equal to the points/div setting

of the oscilloscope.

Version 1.2 update:  When the DocuWave view is not zoomed,

the horizontal scale factor is rounded to a value that will

reflect the value shown on the scope provided the scope is

displaying the entire waveform.  This feature may be disabled

by specifying the /n option on the DocuWave command line.

Application-specific Notes:

=============================================================

--TDS400/A, TDS500/A, TDS600/A, TDS700/A and TDS800 Series

1.  The addition of longer record lengths and Instavue to

some TDS series instruments requires longer acquisition

times for data acquisition and restoration of the

instrument settings.

In order to minimize communication line time-out

problems, it is recommended that DocuWave be set

for 30 seconds (minimum).  For very long record lengths

higher time-out settings may be required.

For record lengths longer than 100,000 points DocuWave

displays the total record in 10 graticules.  At slower

sweep speeds the waveform displayed on the DocuWave screen

will be distorted by a factor determined by the scope

output.  At these longer record lengths the zoom mode is

unable to expand the record so that it may be viewed on

screen.

DocuWave was designed to display the entire waveform in

a 10 graticule display area and so the DocuWave time per

div reflects this. Because record lengths are various sizes

and are displayed at values which differ from 10 divisions

in some cases, DocuWave's vertical and horizontal display

settings will differ from the oscilloscope settings.  This

also applieds if the oscilloscope is placed in the

"Fit To Screen" mode.

--THM 500 Series & Derivatives (ET2020, AVL865, etc.)

1.  If the TekMeter (THM 500 Series & Derivatives) is

set to either SINGLE or NORMAL triggering mode, and

trigger level is set such that the input signal does

not cross the trigger level, and the DocuWave ACQUIRE

key is clicked, the TekMeter will appear to lock up.

After the user-selected communication line time-out

has expired, DocuWave will display a time-out error

panel, and begin a retry.  This will repeat "forever"

without success.  All TekMeter buttons will appear

to be inoperative.

The recommended work-around for this is to click

CANCEL on the DocuWave screen, and either press the

TekMeter Power On/Off button (it will take several

seconds to power off under these conditions), or

remove the TekMeter battery pack.  Then, power up

the TekMeter and adjust trigger threshold so that

the input signal triggers properly.  DocuWave may

be restarted and under these conditions, ACQUIRE

should produce satisfactory results.  IMPORTANT

NOTE:  Any previously-stored settings, screens or

waveforms will be lost if such an event occurs.

2.  DocuWave is able to ACQUIRE a live waveform from a

channel, even if it has been turned OFF at the

TekMeter.  The TekMeter Waveform On/Off function

affects the TekMeter display, only.

3.  DocuWave times out when live waveforms of 10 seconds

per division and slower are ACQUIRED.  Since the

maximum time-out delay selectable under DocuWave is

100 seconds, all acquisitions which take longer than

this will time out.  The work-around is to acquire

the slow (>= 10 s/div) waveform, save it to a reference

using the STO/REC menu selection, STO WAVE.  Then use

DocuWave to retrieve the stored waveform as a REF.

4.  DocuWave can send saved waveforms to the TekMeter only

if the destination reference location already has a

waveform in it.  If the reference waveform is visible

when sending a waveform from DocuWave, the new contents

can be viewed by turning the waveform off and then back

on from the TekMeter front panel.

--THS700 & THS7DMM Series & Derivatives

1.  DocuWave treats the THS700 series like a generic

TDS300 series scope.

Readout of stored waveforms is accomplished by using

Ref A and Ref B.  To obtain the remaining 8 stored

waveforms, you must use the front panel controls of

the THS700 series to map the stored waveforms to

Ref A and Ref B.

2.  To display the DMM logger waveform, requires selecting

the THS7DMM driver.  When the THS7DMM driver is selected only

the DMM logger waveforms can be accessed by DocuWave.

To again access the scope waveforms, requires selection of

the THS710 or THS720 instrument driver.

3.  The THS7DMM driver displays the DMM logger trace on DocuWave's

display as a screen size of 8 X 10 divisions, while the

THSDMM displays the logger trace on a screen size of 8 X 8

divisions.  This results in the time/div differences between

the THS700 series and DocuWave screens.

4.  The THS7DMM driver displays the DMM logger trace on DocuWave's

display with T zero on the left and T max on the right.  This

is the reverse of their positions on the instrument.  (THS7DMM

puts T zero on the right hand end of the trace and T max is on

the left.)

--------------------- End of document ------------------------------