"GRABBER"

Waveform Transfer Software

For the 2201/2211/2214

Digital Storage Oscilloscope

VERSION 2.0

TEKTRONIX UK LTD - HODDESDON

CONTENTS

Chapter 1.0

INTRODUCTION..................................

1.1  Files on GRABBER Distribution Disk

1.2  Hardware and System Software Requirements

1.3  Oscilloscope Parameter Dip Switch Settings

1.4  Installation Requirements

Chapter 2.0

HOW GRABBER WORKS.......................

2.1  GRABBER Operation Modes

Chapter 3.0

RUNNING GRABBER.........................

3.1  GRABBER Command Syntax ...............13

3.1.1  Arguments - n and PORT

3.1.2  Argument  - path\filename

3.1.3  Command Use Examples

3.2  Operating Process

3.2.1  Printer Selection

3.2.2  Sample Waveform Capture Session

3.3  Archiving Captured Waveform Data

3.4  Printing Outside GRABBER

Chapter 4.0

TECHNICAL SPECIFICATIONS................

4.1  HP-GL Emulation

4.2  Files

4.2.1 File Naming Convention

4.3  Communications

4.4  Graphics Display

4.4.1 GRABBER Screen Layout

4.5  Hard Copy Output

4.5.1 Supported Printer Formats

4.5.2 The Printer Driver Program: PRD.EXE

Chapter 5.0

FILE CONVERSION PROGRAM; CONVERT.EXE.....................

5.1  Files Used and Generated by CONVERT.EXE

5.2  USING CONVERT

5.3  Case of 2201 and 2214 Oscilloscopes

Chapter 6.0

ERROR MESSAGES.......................

Appendix A

FILE FORMATS.........................

A.1  GRABBER Waveform Binary Data Files

A.2  TEK Preamble File

A.3  TEK Curve Data

A.3.1 BINARY Format

A.3.2 HEX Format

A.3.3 ASCII Format

A.4  ADIF Format

Appendix B

CABLING REQUIREMENTS..................

Chapter 1.0

INTRODUCTION

The GRABBER software package is for use with the 2201/2211/2214

digital storage oscilloscopes. It emulates a subset of the HP-GL

plotter command language and is used to:

*  Transfer waveforms displayed on the oscilloscope to a

personal computer

*  Store captured waveforms on disk as HP-GL compatible data

*  Display captured waveforms on the computer screen

*  Convert the HP-GL plot data to binary and other formats,

allowing easy transfer to statistical analysis packages etc.

*  Print the captured waveform on a wide selection of

printer/plotters

Capture and transfer of data to the computer is initiated by

pressing the Print/Plot switch on the oscilloscope.

Grabber requires an IBM-PC XT/AT, or a fully compatible

equivalent personal computer.

A wide range of display adapters are supported, including CGA,

EGA, VGA and Hercules. GRABBER automatically detects the graphics

environment, and selects the highest resolution graphics mode

available.

Hard copy plots can be obtained on many printers, using either

parallel or serial interfaces. The printer driver software is a

stand-alone executable program which is called by GRABBER; it can

also be used outside the GRABBER environment to generate hard

copy plots of existing waveform files.

Waveform  data can be converted to ADIF or ASCII, HEX and  BINARY

curve formats that comply with Tektronix Standard Codes &  Format

(062-1780-01  Rev.  C), by means of a conversion program  in  the

package.  Once  converted,  waveform data can  be  downloaded  to

another Tek instrument or to software packages such as SPD (Tek's

Signal  Processing  and  Display package)  etc.  for  statistical

evaluation or analysis.

1.1 Files On the GRABBER Distribution Disk

The GRABBER distribution disk contains the following files:

GRABBER.EXE    Waveform transfer program

PRD.EXE        Printer Driver file used by GRABBER

CONVERT.EXE    Waveform data format Conversion utility

SETCLK14.EXE   Program to set the Real Time Clock on the 2214

oscilloscope. Refer to your instrument manual for

further details.

README.DOC     The GRABBER manual, as a disk file

PLOT.11        An example HP-GL file from the 2211 oscilloscope

1.2 Hardware and System Software Requirement

Operation of GRABBER requires the following

hardware and system software:

*  Tektronix 2201 (with option 12), 2211, or 2214  oscilloscope

*  An IBM PC/XT/AT or fully compatible personal computer with

serial port (COM1 or COM2)

*  PC-DOS (MS-DOS) version 2.0 or later

*  One of the following graphics adapters :

Colour Graphics Adapter (CGA)

Enhanced Graphics Adapter (EGA)

Video Graphics Array (VGA)

Hercules Graphics Adapter

*  Memory requirement: 360K minimum

*  A three-wire cross-over cable (Tx,Rx and Ground) is required

for serial communication between the oscilloscope and the

computer.

This cable is required only when GRABBER is used to capture

data from the oscilloscope.

Please refer to Appendix B for the pin signal assignments

and cable wiring information.

1.3  Oscilloscope Parameter Dip Switch Settings

For GRABBER to work properly, the parameter dip switch settings

on the oscilloscope must be set correctly.  The switch settings

must be as shown below.  On the 2201 and 2211, the setting of S2

is optional, depending on whether or not you wish to display and

transfer the oscilloscope's graticule.

Note: The 9600 baud rate settings shown here, match the default

rate on GRABBER. See section 3.1.1 on how to use other baud rates

with GRABBER.

2201 & 2211 oscilloscopes:

SWITCH & SETTING    ACTION

S1(ON)    Enables software handshaking (XON/XOFF)

S2(ON)    Turns on the graticule (optional)

S3(OFF)   S3 and S4 in combination -

S4(OFF)   Set the communications transfer rate to 9600 baud

S5(ON)    S5 and S6 in combination -

S6(ON)    Select the HPGL4 data output format

2214 oscilloscope:

SWITCH & SETTING    ACTION

S1(ON)    Enables software handshaking (XON/XOFF)

S2(OFF)   S2 and S3 in combination -

S3(OFF)   Set the communications transfer rate to 9600 baud

S4(OFF)   S4 to S7 (inclusive) -

S5(ON)    Select the 4k  HPGL4 data output format

S6(OFF)

S7(OFF)

1.5 Installation Requirements

GRABBER requires no installation if it is to be run from a floppy

disk.  For a hard disk, installation is very simple;  all that is

required is that the GRABBER program files are copied to the disk

from which they are to be run. This can be done as follows:

Before you start, ensure that you are in the root directory of

the required disk.

STEP 1 :

At the MS-DOS prompt, create a new directory called GRABBER

using the MS-DOS MakeDirectory command; you can do this by

typing :

MD GRABBER

STEP 2 :

Use the Change Directory command to move to the new

directory  you have just created; do this by typing :

CD GRABBER

STEP 3 :

Insert the GRABBER software disk into drive A. Copy the

files from the GRABBER software disk to the GRABBER

directory, by typing:

COPY A:*.* C:

STEP 4 :

When all the files are copied, check that all the files from

the distribution disk were copied correctly.

Chapter 2.0

HOW GRABBER WORKS

GRABBER software can be operated in two modes:

*  CAPTURE MODE, where it captures waveform data from the

oscilloscope through the RS232 COM port on the PC, saves the

data on disk, and displays it on the computer screen.

*  DISPLAY MODE, where it reads saved waveform data from a disk

file and displays it on the computer screen.

In both modes, you have the option of generating a printed copy

of the plot.

How to start GRABBER in the required mode is explained in Section

3.1.1.

2.1  GRABBER Operation Mode

When in CAPTURE MODE, GRABBER receives waveform data from the

oscilloscope in HP-GL  (Hewlett-Packard Graphic Language) format

and saves this data in a disk file, named PLOT.n  The software

then reads the PLOT.n file and displays the data it contains on

the computer screen, as a waveform plot.

During this process, GRABBER also decodes the HP-GL data into

appropriate waveform binary data and instrument state

information, and saves the decoded data in a set of disk files.

There can be up to 4 binary waveform data files for each capture,

depending on which channel traces are ON or OFF. There is  always

a single instrument state file, called INSTAt.n.

On the 2201 and 2211 oscilloscopes, the binary files are ACQ1.n,

ACQ2.n for the acquisition channel traces, and REF1.n, REF2.n for

the reference traces.

In the case of the 2214 oscilloscope, with four acquisition

channels and no reference, the REF1 and REF2 files are replaced

by the ACQ3 and ACQ4 files. For the sake of brevity however the

terms REF1 and REF2 are used throughout this manual where in the

case of the 2214 oscilloscope, they refer to ACQ3 and ACQ4

respectively.

In DISPLAY MODE, no communication with the oscilloscope occurs; the

waveform data is read directly from a previously saved disk file

PLOT.n which contains HP-GL plot data. The conversion to the

requisite binary files takes place just as in CAPTURE MODE.

In all the above filenames, the variable n is a decimal number

from 0 to 999, which uniquely identifies the waveform.  Thus,

every file in the set of files generated for a given waveform

display is assigned the same numeric suffix by GRABBER.

Chapter 3.0

RUNNING GRABBER

To run GRABBER, you issue a command at the MS-DOS prompt.

It has two forms which are explained below.

NOTE: Onscreen HELP regarding the use of GRABBER can be obtained

by typing:

GRABBER ?

3.1 GRABBER Command Syntax

The GRABBER command line syntax is:

GRABBER [n] [PORT]

or

GRABBER [pathname\filename]

The n and the PORT arguments together with the Port Baud Rate

selection are explained in section 3.1.1.

The pathname\filename argument has a special use and cannot be

used in combination with either of the other arguments.  It is

explained in section 3.1.2.

3.1.1 Command Arguments - n and PORT

These arguments control the following operational conditions:

*  In CAPTURE MODE:

- the suffix which will identify the files to be

generated

- the communications port on the computer

*  In DISPLAY MODE:

- the previously saved waveform PLOT.n file which is to

be read, for display and possibly printing

The argument n is an integer from 0 to 999, denoting a plot file

number.  This argument must be given on the command line

regardless of the GRABBER operation mode. n is used as suffix to

PLOT.n and all other output files (ACQ1.n, ACQ2.n, REF1.n,

REF2.n, INSTAT.n).

The argument  PORT designates the communications port on the

computer to which the oscilloscope is connected; it can be set to

either COM1 or COM2. GRABBER uses this port to receive data from

the oscilloscope.

The PORT argument controls the mode in which GRABBER runs. If

this argument is left out, GRABBER will start up in DISPLAY MODE.

CAPTURE MODE cannot be entered directly from DISPLAY MODE or vice

versa.

Note on Communication Baud Rate:

Grabber is designed to communicate with the Oscilloscope at a

default speed of 9600 bits/sec. If however the serial port on

your PC does not operate satisfactorily at this rate, you can run

GRABBER with a slower baud rate settings of 4800 or 1200. To do

this, simply type the required value as a third parameter after

the PORT argument. An example is given in section 3.1.3.

3.1.2  Command Argument - path\filename

The path\filename argument identifies an existing file containing

HP-GL compatible data. It allows the display of a previously

stored waveform.

The filename does not have to comply with the PLOT.n form.

In this argument:

path is any valid MS-DOS path specification

filename is any valid MS-DOS filename.

When invoked with this argument, GRABBER merely reads and

displays on the computer screen the specified HP-GL plot file.

No binary or instrument state files are generated.

This version of the GRABBER command is normally used for viewing

archived files without generating any binary or instrument state

files. It does not allow a printout to be obtained.

3.1.3 Command Use Examples

GRABBER 8 <ENTER>

-  Reads the previously saved HP-GL file, PLOT.8 from disk

-  Displays the stored waveform data it contains on the

computer screen

-  Decodes it into the binary files ACQ1.8, ACQ2.8, REF1.8,

REF2.8, and INSTAT.8

GRABBER 15 COM1 <ENTER>

-  Acquires  HP-GL data from the oscilloscope through the

computer's COM1 serial port (at default 9600 baud)

-  Stores this data on disk in PLOT.15 (overwriting the

existing contents)

-  Displays the waveform data on the computer screen

-  Decodes the HP-GL data into the binary files  ACQ1.15,

ACQ2.15, REF1.15, REF2.15, and INSTAT.15

GRABBER 0 COM2 4800 <ENTER>

-  Acquires HP-GL data from the oscilloscope, using port COM2

of the PC as the receiving port. (The baud rate for port

COM2 is set to 4800 baud)

-  Stores this data on disk in PLOT.0 (overwriting the existing

contents)

-  Displays the waveform data on the computer screen

-  Decodes the HP-GL data into the binary files ACQ1.0, ACQ2.0,

REF1.0, REF2.0, and INSTAT.0

GRABBER filename <ENTER>

-  Reads HP-GL data from the disk file (identified  by the

filename argument)

-  Displays the waveform data on the computer screen

-  Does NOT decode the HP-GL data from the input file into

binary files

3.2 Operating Process

After invocation, GRABBER displays an Introduction logo, then

waits for you to press a key. The key press takes you to the main

screen.

If you wish, you can abort the logo drawing process at any time,

by pressing any key.  This facility is useful for lower-

performance PCs, where the logo drawing process can be time-

consuming.

After the logo is completed and a key is pressed, GRABBER

displays the main screen, the basic layout of which is described

in section 4.4.1.

If GRABBER is invoked in CAPTURE MODE, before the appearance of

the main screen prompt, the user is prompted to press the

PRINT/PLOT switch on the oscilloscope, thus  initiating the

transfer of data. The main screen prompt appears after the data

transfer is  completed and the waveform displayed.

Once the waveform plot is displayed, GRABBER waits for you to

press a key.  Valid key options are shown in the prompt list (in

the lower bar of the screen frame).

The prompt list shows the actions that GRABBER can perform next.

The choice of actions are :

*  Quit - press the Q key, and you will be returned to the MS-

DOS command prompt

*  Get plot - press the G key; you will be prompted for a new

plot number to acquire from the oscilloscope or read from a

disk file (depending on the mode in which GRABBER was

invoked).

*  Printer - Press P to display the printer selection screen

(explained in section 3.2.1)

For the G option, the new plot number can be specified in the

following ways:

-  By typing an integer value in the range 0-999 (to select the

next PLOT.n file), then pressing the RETURN key

-  By immediately pressing the RETURN key, to display a list of

all PLOT.n files in your directory, then typing the required

number to select the new PLOT.n file

-  By pressing the ESCAPE key to return to the main screen

If the list of PLOT.n files is too long to fit on a single screen

page, then more files can be displayed by pressing any key.  This

action can be repeated until all the files in the directory have

been listed, or you can terminate the file display by pressing

the ESCAPE key.

NOTE: Only the PLOT file extension numbers are displayed in the

directory listing.

3.2.1  Printer Selection

If you select the P option, you will be presented with a menu of

available printer drivers.  Selecting a driver sends the plot

data shown on the waveform plot display to the selected printer

in an appropriate format. Pressing the ESCAPE key returns you to

the main screen.

To select a printer, you type the number associated with it, then

press the RETURN key.  For example, if you wish to print on an

Epson dot matrix printer you could type 4;  the waveform would be

printed as a half-size plot in high-speed graphics mode.

When you press the RETURN key, GRABBER generates the hard copy

output; this takes a little time.

NOTE: It is essential that you ensure that the printer is plugged

into the computer, before starting to print.  Failure to do  this

will  cause the software to enter and remain in a  waiting  state

until  a printer is connected, just as is the case if the  MS-DOS

"COPY   filename   port"   command  is  used   under   the   same

circumstances.

Please take care in selecting your printer.  If you send output

to the wrong type of printer, the printer/plotter will behave

erratically.

Once printing has commenced, it can be aborted by pressing the

ESCAPE key on the keyboard or turning the printer OFF at the

power switch. Note, however, that in the former case, some print

data is left in the printer's receive buffer and printing will

continue until the buffer is empty.

When printing is aborted, a message is displayed in the top

dialog box of the screen frame. Normal operation can be resumed

at the main screen prompt.

3.2.2 Sample Waveform Capture Session

The following example demonstrates GRABBER being used to read

data  from an oscilloscope into a plot file called PLOT.1, the

plot data being displayed and decoded into the requisite binary

curve data files (ACQ1.1, ACQ2.1) and the instrument state file

(INSTAT.1).

STEP 1 :

The RS232 port of the oscilloscope is connected to the COM1

port of a computer, using the cable specified in Appendix B.

STEP 2 :

The oscilloscope is turned on, and the controls are adjusted

to the following settings :

VERTICAL MODE - BOTH, ALT

VERTICAL POSITION CONTROLS - MIDRANGE

CH1 VOLTS/DIV - 10mV

CH2 VOLTS/DIV - 20mV

AC-GND-DC (CH1 & CH2) - AC

SEC/DIV - 0.2mS

HORIZONTAL MAG - X1

TRIGGER MODE - P-P AUTO

TRIGGER SOURCE - VERT MODE

STORE/NON-STORE - STORE

SAVE/CONTINUE - CONTINUE

Parameter Dip Switches (2201 & 2211 scopes):

S1(ON), S2(ON), S3(OFF), S4(OFF), S5(ON), S6(ON)

Parameter Dip Switches (2214 oscilloscope):

S1(ON), S2(OFF), S3(OFF), S4(OFF), S5(ON), S6(OFF),

S7(OFF), S8(x)

STEP 3 :

Two X10 probes from the CH1 and CH2 inputs are connected to

the oscilloscope's PROBE ADJUST terminal.

STEP 4 :

The GRABBER directory of the computer is entered, and the

following command is issued to start GRABBER  :

GRABBER 1 COM1 <ENTER>

As a result, the program starts up in CAPTURE MODE, and

displays the introductory logo, with the message

Press Any Key To Continue

STEP 5 :

The <ENTER> key on the computer keyboard is pressed; the

program responds with the message:

Waiting For Data -Press Print/Plot Switch or 'Q' To Abort

The program will not continue until the Print/Plot switch

has been pressed on the oscilloscope. Alternatively,

pressing the Q key on the computer keyboard would result in

the program aborting.

NOTE: Under certain circumstances, the message:

Error 9 : Already Plotting - Data Lost!

can appear.  The usual cause is a previous plot which was

aborted from GRABBER during transfer; the oscilloscope is

still sending data from this previous session.

To recover from this condition, completely abort the plot by

briefly pressing the Print/Plot switch then press it again

and  hold in for three seconds.(In the case of  2214,  press

Plot  button once until the Plot light goes  out).  Finally,

re-issue the GRABBER command, as specified in STEP 4.

STEP 6 :

The Print/Plot switch on the oscilloscope is pressed; the

program  responds with the message:

Reading Plot Data - Hit 'Q' To Abort

HP-GL Data is now being transferred into the file PLOT.1.

The data transfer can be aborted by pressing the Q key on

the keyboard or pressing the oscilloscopes's Print/Plot

switch for three seconds (in the case of 2214, until the

Plot light goes out). GRABBER will then display and decode

any data that has been captured so far.

After a successful transfer, the data displayed will look

the same as the oscilloscope trace display at the time the

Print/Plot switch was pressed.

STEP 7 :

From this point, the lower prompt bar gives a choice of

actions:

-  The Q key on the computer keyboard can be used to exit

from GRABBER.

-  The G key can be pressed to continue GRABBER operation

with another plot; a further prompt would solicit a new

plot number, and subsequently initiate another waveform

transfer.

-  The P key can be used to obtain a printed copy of the

saved plot file.

3.3  Archiving Captured Waveform Data

If you wish to retain captured waveform plot data for future

reference, you can do so by archiving the PLOT.n file.  You do

not need to keep the ACQx.n, the REFx.n, or the INSTAT.n files,

as they can be regenerated from the relevant PLOT.n file simply

by using this file as input to GRABBER in DISPLAY MODE.

3.4 Printing Outside GRABBER

To print a PLOT file at a printer that is not attached (or

networked) to your PC, you can use  the stand-alone printer

driver program PRD.EXE, which is supplied on the GRABBER

distribution disk.

Obtaining printed copy requires two things :

-  A copy of the PRD.EXE program file on the PC attached to the

printer

-  Copies of the waveform binary files (ACQx.n and REFx.n) and

the instrument state file (INSTAT.n) for the saved waveform

Note: PRD.EXE cannot print directly from a PLOT.n file.

The command line argument to run PRD is as follows:

PRD <File Extension Number> <Printer Format>

You will find information on the variables, etc, in section 4.5.2

of this manual.

Chapter 4.0

TECHNICAL SPECIFICATION

4.1  HP-GL Emulation

GRABBER emulates a subset of HP-GL plotter command language and

fully supports the HP-GL 4 pen plotter output on the

2201/2211/2214 oscilloscopes.

The following HP-GL commands are supported by the GRABBER

software :

PD                            Pen Down

PU                            Pen Up

PD x,y                    Move with Pen Down

PU x,y                    Move with Pen Up

DF                            Set Defaults

SC Xmin,Xmax,Ymin,Ymax    Scale

SP n                      Select Pen

PA x,y                    Plot Absolute

SR width,height           Set Relative character size

LB Text String            Label

4.2  Files

GRABBER receives waveform data from the scope in HP-GL format,

then displays the waveform on the PC's screen, and saves the data

on disk as a PLOT.n file.

The HP-GL data from the scope is also decoded to generate an

instrument state file (INSTAT.n) and up to four raw data files,

corresponding to acquisition channels 1 & 2 (files ACQ1.n,

ACQ2.n) and reference channels 1 & 2 (files REF1.n, REF2.n). The

INSTAT.n file is for internal use by GRABBER programs only.

NOTE: For the 4-channel 2214 scope, reference waveforms are

replaced by the additional acquisition channels ACQ3 and ACQ4,

giving a total of four ACQx.n files.

4.2.1  File Naming Convention

All files generated by GRABBER have a numeric extension (0-999

decimal). All files corresponding to a particular waveform screen

capture have a common extension number.

4.3 Communications

GRABBER uses software handshaking (XON/XOFF) to communicate with

the oscilloscope.

Data Transmission commences when the scope's Print/Plot switch is

depressed.

GRABBER operates at a default baud rate of 9600 bits/sec. Slower

speeds (4800 or 1200 baud) can also be used if your PC's serial

port does not operate satisfactorily at 9600 baud. refer to

section 3.1.1 for more details.

The serial data format used is fixed and as follows:

1 start bit

8 data bits

1 stop bit

No parity

On start up, GRABBER automatically configures the selected COM

port to the above settings, and restores to original settings

upon termination.

4.4 Graphics Display

GRABBER automatically detects the display environment and selects

the highest resolution graphics mode available.

Most widely available graphics adapters are supported, including:

- EGA

- VGA

- CGA

- Hercules

The drivers for these adapters are built in to GRABBER; no

installation is necessary.

4.4.1 GRABBER Screen Layout

There are 3 distinct areas on the GRABBER main screen:

*  The TOP rectangular area, which is used for PLOT file names

and warning messages

*  The BOTTOM rectangular area, which is used for prompts and

general information

*  The CENTRE square, which represents the CRT screen of the

oscilloscope; all plots are displayed here. It is also used

to display the printer options menu and file listings

The HP-GL plot coordinates are mapped into the CENTRE square,

using scaling factors defined in the HP-GL file.

4.5 Hard Copy Output

GRABBER provides facilities to produce hard copy output of screen

plots.

All data is sent to the standard list device on the PC, normally

the parallel port (LPT1). This is used to drive a  parallel

printer/plotter.

Using A Serial Printer:

To use a serial printer, the LPT1 port must be rerouted to a

serial port (COM1 or COM2) of the computer using the MODE

command in MS-DOS. This must be done before running GRABBER.

For example if your serial printer is connected to the COM2

port of the computer, first ensure that the communication

parameters (e.g. baud rate, data bits, parity etc.) on COM2

and the printer are matched together (Note: MODE command can

be used to set the COM2 serial parameters), then issue the

following MODE command:

MODE  LPT1: = COM2:

Check that this rerouting operation has worked successfully

by dumping a small text file to your serial printer using

the following MS-DOS COPY command:

COPY textfile LPT1:

NOTE: for an HP-GL plotter, the textfile must be in HP-GL

format.

All data sent to LPT1 should now be rerouted to COM2 which

is connected to your printer. You can now go ahead and use

GRABBER or the PRD.EXE program to obtain hardcopy plots on

your serial printer/plotter.

For details of the MODE command, please consult your MS-DOS

handbook.

4.5.1 Supported Printer Formats

GRABBER allows you to select the printer type and plot size from

a menu of available printer drivers.

Four printer standard formats are supported which are used by

majority of printer manufacturers. These include:

* HPGL plotter format

-- 1 pen plot

-- 4 pen plot

* EPSON Graphics mode

-- EPS7 "Esc L" mode

-- EPS8 "Esc Y" mode: gives lower resolution output than

EPS7, but is faster

* HP LaserJet format

* HP ThinkJet format

4.5.2 GRABBER Printer Driver Program: PRD.EXE

The Printer Driver program PRD is a standalone, executable

program that is called by GRABBER to carry out hardcopy

operations. PRD can be used outside the GRABBER environment to

obtain a hard copy print of the plot data.

PRD uses the binary raw waveform data files (ACQ1.n, ACQ2.n,

REF1.n, REF2.n) and the corresponding instrument state file

(INSTAT.n) to produce hardcopy plot.

If no waveform appears on the plot, then it is likely that PRD

has not found any of the binary waveform files or the INSTAT file

in the current directory.

To obtain onscreen help regarding PRD, type:

PRD ?

The command line argument to run PRD is as follows:

PRD <file extension number> <printer format>

The file extension number and printer format must be specified

using the following values:

file extension number

- an integer in the range of 0 to 999

printer format

-  an integer value which specifies the printer type and

plot size as shown in the table below:

Value        Printer type/plot size

0               HP-GL 4 pen plotter

1               HP-GL 1 pen plotter

2               EPS7  Epson Dot Matrix half size (Hi-Res)

3               EPS7  Epson Dot Matrix full size (Hi-Res)

4               EPS8  Epson Dot Matrix fast half size

5               EPS8  Epson Dot Matrix fast full size

6               HP LASERJET half size

7               HP LASERJET full size

8               HP THINKJET half size

9               HP THINKJET full size

Chapter 5.0

FILE CONVERSION PROGRAM; CONVERT.EXE

CONVERT is a file format conversion utility that allows users to

convert GRABBER-generated waveform data files (ACQx.n, REFx.n)

into ADIF or ASCII, HEX and BINARY formats compatible with

Tektronix Standard Codes & Formats (062-1780-01 Rev.C). It uses

the GRABBER's (INSTAT.n) file in the process to generate

readable waveform preamble information.

5.1 Files Used and Generated by CONVERT

CONVERT takes a single waveform binary file (ACQx.n or REFx.n) as

input and uses the corresponding instrument state file (INSTAT.n)

to generate one or two new files depending on the output format

chosen.

When ADIF format is requested, the converted waveform preamble

and curve data are placed into a single file  called ADIFWAV.n.

In the case of other formats, the preamble and curve data are

generated into two separate files called TEKPRE.n and TEKCRV.n.

TEKCRV.n is the Tek curve data file and can be generated in

ASCII, HEX or BINARY format.  The output format is selected by

the user in response to a prompt from CONVERT.EXE.

TEKPRE.n  is a readable ASCII file containing  waveform  preamble

information e.g. Number of points per curve, scale factors,

horizontal increment, scaling unit, data encoding, etc.

ADIFWAV.n is an ascii file containing the waveform preamble and

curve data in the form of Tektronix ADIF format.

5.2 Using CONVERT

To obtain onscreen help regarding CONVERT, type:

CONVERT ?

The command to run CONVERT is :

CONVERT Filename

The filename argument denotes the name of the (ACQx.n or REFx.n)

file to be converted.

The filename must have a valid numeric extension number between 0

to 999.  The extension is required because it is extracted from

the filename and used to read and convert the corresponding

INSTAT.n file.

If either of the input files is not found, the program aborts; no

conversion is performed.

Once CONVERT.EXE has found the specified waveform data file and the

corresponding INSTAT file, it prompts the user for the type of

output format required.

The format of the output files depend on the type of conversion

chosen. Refer to the previous section for a description of the

various output files.

You can override the default output file names by specifying new

names when prompted by the program.

NOTE: The INSTAT file which is selected by CONVERT.EXE relates

only to the current acquisition channel waveform (ACQx).

If a file other than the current acquisition waveform data file

is to be converted (for example, REF2.n), the INSTAT file may not

truly represent the state of this curve and the generated

preamble in TEKPRE.n or  ADIFWAV.n files may need to be edited

manually to reflect the correct state of the waveform.

5.3 Note for 2201 and 2214 Oscilloscopes

In the case of 2201 and 2214 Oscilloscopes, the full waveform

preamble information is not available, so parts of the generated

TEKPRE.n file and the preamble section of the ADIFWAV.n file are

defaulted to a predetermined value. These include:

- vertical attenuation - set to 1 volts/div

- time base - set to 1 sec/div

The default values can be modified to your own requirements with

a suitable text editor.

Details of the file formats are given in Appendix A.

Chapter 6.0

ERROR MESSAGES

When an error condition arises, the GRABBER program terminates

and displays one of the following messages.

Error 1 : Graphics System Error.

Indicates that the graphics adapter of the PC is either not

recognized by GRABBER or not operational.

Error 2 : Could Not Open Port COMx.

Hardware error - the COM port specified is either not operational

or not available.

Try running GRABBER with an alternative COM port argument.

Error 3 : Software Handshake Failed - Data Lost !

Occurs when either:

-  software handshaking is not operating or

-  the oscilloscope is transmitting at a different baud rate to

GRABBER

Ensure that the serial interface dip switches on the oscilloscope

are set for software handshaking and 9600 baud transmission, as

explained in section 1.3.

Error 4 : Could Not Open File.

Indicates a non-existent file or disk error when opening a file.

Error 5 : Disk Error When Writing To File.

A disk problem has occurred while writing to a file (for example,

the disk is full).

Error 6 : Too Many Command Line Parameters !

More parameters have been specified than are allowed in the

command line syntax used by GRABBER.

Error 7 : Syntax Error - n Is Not A Valid Plot Number.

An invalid plot number was specified in the GRABBER command line

argument.

A valid plot number is an unsigned integer in the range 0 to 999.

Error 8 : Could Not Find File Or Port.

The file or COM port specified does not exist.

Error 9 : Already Plotting - Data Lost !

Data was transmitted before GRABBER was initialized to capture

it.

To remedy, press the Print/Plot switch on the oscilloscope once,

then press the switch again and hold for 3 seconds; finally,

rerun GRABBER. (In the case of 2214, press Plot button once until

the Plot light goes out).

Error 10: Unrecognized HP-GL Format - Aborting.

The transferred data was not in HP-GL format.

If GRABBER was receiving data from the oscilloscope, ensure that

the serial interface dip switches are set for HPGL4 format, as

explained in section 1.3.

Error 11: Unrecognized HP-GL Command - Aborting.

The plot file contained HP-GL commands unrecognized by GRABBER.

Error 12: Could not find Printer Driver file (PRD.EXE).

GRABBER cannot find PRD.EXE in the current directory.

Error 13: Not enough memory to run Printer Driver program.

Error 14: General Error in executing Printer Driver program.

Appendix A

FILE FORMATS

A.1  GRABBER Waveform Binary Data File

Up to four binary waveform data files are generated, depending on

the status of traces within the oscilloscope; that is, if they

are ON or OFF.

These files are named ACQ1.n, ACQ2.n, REF1.n and REF2.n (ACQ3.n

and ACQ4.n in the case of the 2214 oscilloscope). They contain

raw curve data.

Each file consists of 4000 bytes (8-bit unsigned integer) of data

(2000 bytes in the case of 2201 oscilloscope).  Data values

correspond to the Y deflection values on the oscilloscope.

NOTE: Lower graticule corresponds to a value of 25 and there are

25 points per division.

A.2  TEK Preamble File

The Waveform preamble indicates the waveform attributes, such as

number of points per waveform, scale factor, offset, horizontal

increment, scaling units and data encoding. This file which is in

ASCII form, is produced from the GRABBER INSTAT.n file.

A typical preamble data file is as follows:

WFM WFI:"2211, ACQ1, 5V, AC, 0.1ms",

NR.P:4000,PT.O:1024,PT.F:Y,

XMU:1.0E+0,XOF:0,XUN:S,XIN:2.5E-07,YMU:2.0E-01,

YOF:0,YUN:V,

ENC:BIN,BN.F:RP,BYT:1,BIT:8,CRV:CHK;

The file:

-  starts with the header WFM (short for WFMpre), indicating a

waveform preamble file

-  ends with a Carriage Return/Line Feed as terminator

A ";" is added before the terminator, to act as a delimiter

between preamble and curve data if the two are to be concatenated

into one file.

The individual fields, their arguments and relevant description

are given below.

Each field is given in BOLD letters. The minimum field name

appears in capital letters with the remaining in lower case.

The possible arguments for that field are shown in front of it

followed by a description note. The argument <NR1> indicates a

signed integer value. The argument <NR3> indicates a floating

point value.

FIELD        ARGUMENT

WFId:    "ascii string"

The WFI field includes labeling information to help you

remember key features about the waveform. The information

includes the Scope Type, Vertical Mode, Coupling, Volts/Div

and Time/Div. The scaling information is the same as in the

corresponding preamble fields, but is labeled in the

appropriate units.

An "ADD" in this field indicates vertical Add mode, an "INV"

means that Acquisition channel 2 is inverted, an "ADD-

UNKNOWN" indicates unknown scaling factor when adding two

channels with at least one being uncalibrated. Also a ">" in

front of the vertical scaling factor indicates variable

deflection i.e. the acquisition channel is in Uncal mode

(Note: YUNits will be set to DIV and YMUlt to 0.04 in this

mode).

Note that in the case of 2201/14 oscilloscopes since the

full preamble information is not available, only the type of

scope and the channel name appears in this field.

NR.Pts:        <NR1>

This field defines the number of points in the waveform. It

indicates the length of record transferred.

Pt.Off:        <NR1>

Point Offset identifies the trigger position relative to the

first point of the waveform.

PT.Fmt:        Y

Point Format defines how to interpret the curve data. Y is

always used; the data points relate to the Y values on the

Y-T display.

XUNits:        S or CLK

If the argument is S then XINCr value is in seconds. If it

is CLKs (External Clock source), then the scaling is

unknown.

XINcr:         <NR3>

Value gives the time interval between sample points. It is

calculated as:

(sec/div) / (No_of_Points/div).

Example: for a time base of 0.1 milli sec/div and 4000

points across 10 divisions horizontally

XIN = (0.1E-03) / (4000/10) = 2.5E-05

YUNits:        V, DIV

Indicates units associated with YMUlt. V represents Volt

units; DIV represents variable deflection (i.e. uncalibrated

mode)

YMUlt:          <NR3>

This value gives the vertical step size of the digitizer

(volts between adjacent quantization levels). It is

calculated as:

(volts/div) / (points/div)

Example: for a vertical deflection of 5 volts/div and 25

points per vertical division

YMU = 5 / 25 = 2.0E-01

Note: If YUNits = DIV, then YMUlt is unknown and set to 1/25

= 0.04

YOFf:          <NR1>

This is the Y coordinate of ground. Since this information

is not available on the supported oscilloscopes, it is set

zero.

XMUlt:

XOFf:

Analogous to YMUlt and YOFf. They are not used for the

2201/2211/2214 oscilloscopes, and hence are set to 1.0 and

0.0 respectively.

ENCdg:         ASCii, BINary or HEX.

Determines waveform encoding for the TEKCRV.n file.

BN.Fmt:        RP

Binary format is always a Right-justified, Positive binary

integer.

BYT:           <NR1>

Always set to 1 indicating that each data value is contained

in one byte.

BIT:           <NR1>

Each data value consists of 8 bits.

CRVchk:        CHKsm0

CHKsm0 indicates that the last byte of a binary curve is a

checksum. It is the 2's complement of the modulo 256 sum of

the binary count and curve data bytes. It does not include

the header "CURVE %" that precedes the binary count.

A.3  TEK Curve Data

The TEK curve data files have the following components:

*  All TEK curve data files start with the string "CURVE",

followed by a space.

*  Then follows an identifier for the data format, as shown

below:

"%"   for BINARY format.

"#H"  Two character combination for HEX format.

" "  nothing for ASCII format.

*  Then in the case of BINARY & HEX, a curve Data Count value.

*  Next is the Curve Data block.

*  Then, for BINARY & HEX files only, a Checksum value.

*  Then a  ";" for delimiter.

*  Finally a Carriage Return/Line Feed Terminator.

The BINARY, HEX and ASCII format of TEK curve data are explained

in the subsections of A.3. The ADIF ascii format is described in

section A.4.

A.3.1 BINARY Format

With this format, the waveform curve data is in the form of:

CURVE <space> % <Binary Count MSB> <Binary Count LSB> <Data>

<Checksum> ; <Terminator>

Where:

%                    is used as a header character to show the

start of a Binary block.

<Binary Count MSB>

is the most significant byte of the two byte

Binary Count. Binary Count is the length of

the waveform, in bytes, plus the one byte

Checksum.

<Binary Count LSB>

is the least significant byte of the Binary

Count.

<Data>

is made up of 4000 (or 2000 for 2201) data

points. Each data point is a 1 byte (8-bit)

representation of each digitized value.

<Checksum>

is the two's-complement of the modulo 256 sum

of the preceding data bytes and the binary

count. The checksum is used to verify that

all data values have been received correctly.

;               is the block delimiter.

<CR><LF>       Carriage Return Line Feed Terminator.

No delimiter is used between individual data points.

A.3.2 HEX Format

Data Format in HEXadecimal waveform encoding is very similar to

Binary format with the following exceptions:

-  The curve header is "CURVE #H" instead of "CURVE %".

-  Each 8-bit data point is represented by 2 ASCII hexadecimal

characters.

-  The byte count is sent as four successive ASCII hexadecimal

characters, but the value of the byte count itself is

identical to a comparable BINARY curve data.

The Checksum is sent as two successive ASCII hexadecimal

characters.

The format of the HEX file is:

CURVE<space>#H                  (8 bytes header)

<MS 4 bits of MSB Bin Count>      (byte 9)

<LS 4 bits of MSB Bin Count>      (byte 10)

<MS 4 bits of LSB Bin Count>      (byte 11)

<LS 4 bits of LSB Bin Count>      (byte 12)

<1st Point MS 4 bits>             (byte 13)

<1st Point LS 4 bits>             (byte 14)

<2nd Point MS 4 bits>             Etc..

<2nd Point LS 4 bits>

...

...

<nth point MS 4 bits>

<nth point LS 4 bits>

<ChkSum MS 4 bits>

<ChkSum LS 4 bits>

;

<CR>

<LF>

A.3.3 ASCII Format

In ASCII encoding, each data point is represented by a variable

length integer; e.g. a byte value of D6(hex) is represented by an

integer value of 214. Integer values are separated by commas in

the curve data block section.

There is no Data Count value before the <data> block and no

Checksum after it.

In ASCII format, the waveform data is represented as:

CURVE<space>data,data,data,...,data;<Terminator>

Each data point represents a variable length integer.

A semicolon is inserted after the last data point.

<Terminator> is a Carriage Return Line Feed combination.

A.4 ADIF Format

The Tektronix Analog Data Interchange Format (ADIF) is intended

to facilitate the sharing of waveform and other analog

information between software packages and instruments supporting

it. It is flexible, extendible, and can accommodate a wide range

of data structures and formats.

ADIF is block-structured. An ADIF data set generally consists  of

several description blocks and a data block. Each block addresses

a different aspect of data description.

The ADIF file generated by the CONVERT.EXE program represents a

subset of the ADIF standard. It can be imported to a variety of

software packages e.g. SPD (Tek's Signal Processing and Display

package) etc. for statistical evaluation or analysis.

The structure of a typical ADIF file is shown below:

ADIF = <data set name> (STD (VERSION 0.9) |

IDENTIFY (

NAME "Scope's Name" |

NOTE "" |

DATE 1989-08-15 |

TIME 12:54:14

) |

DIM = <x dim name> (

TYPE IMPLICIT |

SCALE <x scale> |

OFFSET <x offset> |

UNITS <x units> |

SIZE <size>

) |

DIM = <y dim name> (

TYPE EXPLICIT |

SCALE <y scale> |

OFFSET <y offset> |

UNITS <y units> |

) |

ENCODE (FORMAT INT8) |

DATA (CURVE (VALUES

<data block> ) )

)

The revision level of the ADIF standard itself and any optional

protocols used are described in the STD block.

<data set name> =   Any appropriate name for data set, consisting

of 12 characters or less. Here it represents

the appropriate vertical channel.

The IDENTIFY block describes the environment and manner in which

the data was obtained. The NOTE field is used in the following

cases: vertical ADD mode, Channel 2 INVert mode and UNKNOWN

scaling factor (i.e. Uncalibrated). For more details on this

refer to WFId section of Appendix A.2 on Tek preamble data file.

The DIMension and ENCODE blocks describe how the data are

physically represented and logically organized. The ENCODE block

is not necessary if the curve data values are in ascii format.

The CONVERT.EXE program therefore does not generate the ENCODE

block.

<x dim name>  =     Any appropriate name (up to 12 characters)

for the x dimension. (no enclosing quotes).

<x scale>     =     XINcr (as in Tek's Codes & Formats WFMpre)

<x offset>    =     XOFf (as in Tek's Codes & Formats WFMpre)

<x units>     =     XUNit (as in Tek's Codes & Formats WFMpre)

<size>        =     NR.PT (as in Tek's Codes & Formats WFMpre)

Number of points in the curve data.

<y dim name>  =     Any appropriate name (up to 12 characters)

for the y dimension. (no enclosing quotes).

<y scale>     =     YMUlt (as in Tek's Codes & Formats WFMpre)

<y offset>    =     YOFf (as in Tek's Codes & Formats WFMpre)

<y units>     =     YUNit (as in Tek's Codes & Formats WFMpre)

For a description of Tek's Codes & Formats WFMpreamble refer to

section A.2 on Tek preamble file TEKPRE.n.

TYPE EXPLICIT: indicates that the values for this dimension are

present in the DATA(CURVE).

TYPE IMPLICIT: indicates that the values for this dimension are

not in DATA(CURVE) but are instead derived from

the location of other coordinate values; e.g. for

a Y-T display  the nth Y deflection value in the

DATA(CURVE) represents the nth point in Time for

the X (Time) axis.

The DATA block, contains the waveform data point values. These

are the unscaled sample point values. Data values are ascii

integers separated by commas. There are <size> number of values

in this block.

Appendix B

CABLING REQUIREMENTS

A three-wire cable (TXD,RXD and Ground) is required for serial

communication between the oscilloscope and the computer.

The cable must have a 9-way female connector at the oscilloscope

end and a 9- or 25-way female connector at the computer end

(depending upon the type of computer serial COM port). The

Transmit (TXD) and Receive (RXD) lines at one end of the cable

must be cross wired to the RXD and TXD at the other.

Assuming the following convention for the Computer's serial

ports:

9-way:   TXD at pin 3  and RXD at pin 2

25-way:  TXD at pin 2  and RXD at pin 3

The wirings are as shown below:

Oscilloscope's port      Computer's 9-way serial port

PIN                 PIN

RXD (in)   2  ---------------- 3   TXD (out)

TXD (out)  3  ---------------- 2   RXD (in)

GROUND     5  ---------------- 5   GROUND

Oscilloscope's port      Computer's 25-way serial port

PIN                 PIN

RXD (in)   2  ---------------- 2   TXD (out)

TXD (out)  3  ---------------- 3   RXD (in)

GROUND     5  ---------------- 7   GROUND